{"text_format":"1","workbook":"xlsx","n_sheets":2,"sheets":["Laser Form","Dropdowns"],"not_rendered":{"charts":0,"images":1,"macros":1,"external_links":0,"defined_names":3,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Laser Form","index":0,"n_rows":191,"n_cols":6,"rows_kept":191,"cut":false,"hidden":false,"merged":[[188,1,188,5],[110,2,110,3],[104,2,104,3],[124,5,131,5],[115,2,115,3],[104,5,117,5],[114,2,114,3],[18,4,18,5],[19,4,19,5],[20,4,20,5],[22,4,22,5],[21,4,21,5],[23,4,23,5],[78,5,81,5],[60,5,62,5],[24,4,24,5],[25,4,25,5],[27,4,27,5],[28,4,28,5],[57,5,59,5],[67,5,77,5],[30,4,30,5],[26,4,26,5],[82,5,83,5],[55,1,55,5],[133,1,133,3],[135,2,135,3],[142,2,142,3],[82,2,82,3],[66,1,66,3],[84,2,84,3],[94,2,94,3],[104,1,117,1],[108,2,108,3],[96,2,96,3],[97,2,97,3],[113,2,113,3],[101,2,101,3],[102,2,102,3],[127,2,127,3],[117,2,117,3],[123,2,123,3],[121,2,121,3],[100,2,100,3],[125,2,125,3],[126,2,126,3],[103,2,103,3],[111,2,111,3],[112,2,112,3],[139,2,139,3],[140,2,140,3],[141,2,141,3],[134,1,136,1],[137,1,142,1],[134,2,134,3],[136,2,136,3],[137,2,137,3],[138,2,138,3],[124,1,131,1],[129,2,129,3],[119,2,119,3],[120,2,120,3],[85,1,85,3],[118,1,123,1],[130,2,130,3],[118,5,123,5],[107,2,107,3],[91,1,91,3],[86,1,89,1],[89,2,89,3],[90,2,90,3],[92,1,103,1],[116,2,116,3],[118,2,118,3],[122,2,122,3],[105,2,105,3],[106,2,106,3],[98,2,98,3],[99,2,99,3],[95,2,95,3],[92,5,103,5],[109,2,109,3],[72,2,72,3],[82,1,83,1],[73,2,73,3],[74,2,74,3],[81,2,81,3],[83,2,83,3],[56,1,56,3],[65,2,65,3],[67,2,67,3],[68,2,68,3],[69,2,69,3],[70,2,70,3],[57,1,59,1],[60,1,62,1],[63,1,64,1],[59,2,59,3],[64,2,64,3],[61,2,61,3],[62,2,62,3],[63,2,63,3],[71,2,71,3],[3,1,3,5],[48,1,48,2],[49,1,49,2],[51,1,51,2],[50,1,50,2],[44,4,44,5],[45,4,45,5],[46,4,46,5],[44,1,44,2],[45,1,45,2],[46,1,46,2],[48,4,48,5],[49,4,49,5],[50,4,50,5],[51,4,51,5],[5,3,5,5],[6,3,6,5],[7,3,7,5],[8,3,8,5],[10,3,10,5],[11,3,11,5],[29,4,29,5],[16,4,16,5],[17,4,17,5],[134,5,136,5],[1,1,1,5],[86,2,86,3],[87,2,87,3],[88,2,88,3],[92,2,92,3],[93,2,93,3],[75,2,75,3],[77,2,77,3],[78,2,78,3],[79,2,79,3],[80,2,80,3],[78,1,81,1],[53,1,53,5],[32,1,32,5],[14,1,14,5],[57,2,57,3],[58,2,58,3],[60,2,60,3],[12,3,12,5],[67,1,77,1],[63,5,64,5],[76,2,76,3],[10,1,11,1],[182,1,182,5],[185,2,185,3],[146,2,146,3],[147,1,147,3],[148,1,151,1],[148,2,148,3],[131,2,131,3],[128,2,128,3],[124,2,124,3],[132,2,132,3],[144,2,144,3],[143,1,146,1],[184,2,184,3],[150,2,150,3],[143,2,143,3],[145,2,145,3],[155,1,156,5],[157,1,157,5],[153,1,153,5],[148,5,151,5],[149,2,149,3],[151,2,151,3],[143,5,146,5],[137,5,142,5]]}
[]
[null,"COMPREHENSIVE LASER FORM"]
[]
[null,"I. General Information:"]
[]
[null,"A. PI Information","PI Name:"]
[null,null,"Department:"]
[null,null,"Email Address:"]
[null,null,"Phone No.:"]
[]
[null,"B. Designated Laser Manager","Name:"]
[null,null,"Email Address:",{"$f":"IF(OR(ISBLANK(D11),ISBLANK(D6),ISBLANK(D8)),\"\",IF(D11=D6,D8,\"\"))","v":""}]
[null,null,"Phone No.:",{"$f":"IF(OR(ISBLANK(D11),ISBLANK(D6),ISBLANK(D9)),\"\",IF(D11=D6,D9,\"\"))","v":""}]
[]
[null,"II. Laser for Registration"]
[]
[null,"Bldg. Name or No.","(e.g. B120 or PSB)",null,{"$f":"IF(D17=\"\",\"Enter Room/Lab No.\",\"\")","v":"Enter Room/Lab No."}]
[null,"Room / Lab. No.","(e.g. R249 or 249)",null,{"$f":"IF(D18=\"\",\"Enter Building Name or No.\",\"\")","v":"Enter Building Name or No."}]
[null,"Laser Serial No.",null,null,{"$f":"IF(D19=\"\",\"Enter Serial No.\",\"\")","v":"Enter Serial No."}]
[null,"Model",null,null,{"$f":"IF(D20=\"\",\"Enter Model\",\"\")","v":"Enter Model"}]
[null,"Manufacturer",null,null,{"$f":"IF(D21=\"\",\"Enter Mfr.\",\"\")","v":"Enter Mfr."}]
[null,"Class","3b or 4",null,{"$f":"IF(D22=\"\",\"Select from Drop-down\",\"\")","v":"Select from Drop-down"}]
[null,"Medium","(Argon, CO2, Nd:YAG, etc.)",null,{"$f":"IF(D23=\"\",\"Enter Medium\",\"\")","v":"Enter Medium"}]
[null,"Type","Continuous, Pulse, Repetitive Pulse",null,{"$f":"IF(D24=\"\",\"Select from Drop-down\",\"\")","v":"Select from Drop-down"}]
[null,"Wavelength","(nm)",null,{"$f":"IF(D25=\"\",\"Enter value\",\"\")","v":"Enter value"}]
[null,"Power Output","(W)",null,"Enter value"]
[null,"Beam Diameter","(mm)",null,"Enter value"]
[null,"Divergence","(mrad)",null,"Enter value"]
[null,"Exposure Duration (T)","(s)",{"$f":"IF(D24=\"Continuous\",10,IF(OR(D24=\"Pulsed\",D24=\"Repetitive Pulse\"),0.25,\"\"))","v":""},{"$f":"\"Default at 10s if CW and 0.25s if Pulsed and Repetitive\"","v":"Default at 10s if CW and 0.25s if Pulsed and Repetitive"}]
[null,"Pulse Duration (t)","(s)",null,{"$f":"IF(D30=\"\",\"Enter value if Pulsed or Repetitive\",\"\")","v":"Enter value if Pulsed or Repetitive"}]
[null,"Pulse repetition frequency","(Hz)",null,{"$f":"IF(D31=\"\",\"Enter value if Repetitive\",\"\")","v":"Enter value if Repetitive"}]
[]
[null,"III. Hazard Analysis"]
[]
[null,"Irradiance, E ","(W/cm2)",{"$f":"IF(ISBLANK(D26),\"\",D26/((PI())*0.35^2))","v":""}]
[null,"No. of Pulses","(N)",{"$f":"IF(AND(D24=\"\", D31=\"\", D30=\"\"),\n   \"\",\n   IF(D24=\"Repetitive Pulse\",\n      IF(OR(D31=\"\", D30=\"\"), \"\", VALUE(D31)*VALUE(D30)),\n      \"n/a\"))","v":""}]
[null,"Energy per pulse","(J)",{"$f":"IF(AND(D24=\"\", D26=\"\", D31=\"\"),\n   \"\",\n   IF(D24=\"Repetitive Pulse\",\n      IF(OR(D26=\"\", D31=\"\", NOT(ISNUMBER(D26)), NOT(ISNUMBER(D31))), \"\", D26/D31),\n      \"n/a\"))","v":""}]
[null,"MPE for Single Pulse Limit (Rule 1)","(J/cm2)",{"$f":"IF(AND(ISBLANK(D24),ISBLANK(D25),ISBLANK(D30),ISBLANK(D31)),\"n/a\",IF(D24=\"Repetitive Pulse\",IF(AND(ISNUMBER(D30),ISNUMBER(D31)),_xlfn.IFS(AND(D25>=180,D25<=260,D30>=0.0000000000001,D30<=0.00000000001),0.0003,AND(D25>=180,D25<=260,D30>=0.00000000001,D30<=0.000000001),0.001,AND(D25>=180,D25<=260,D30>=0.000000001,D30<=10),0.56*D30^0.25,AND(D25>=260,D25<=302.5,D30>=0.0000000000001,D30<=0.00000000001),0.0003,AND(D25>260,D25<=302.5,D30>=0.00000000001,D30<=0.000000001),0.001,AND(D25>=260,D25<=302.5,D30>=0.000000001,D30<=10),0.56*D30^0.25,AND(D25>=302.5,D25<=315,D30>=0.0000000000001,D30<=0.00000000001),0.0003,AND(D25>=302.5,D25<=315,D30>=0.00000000001,D30<=0.000000001),0.001,AND(D25>=302.5,D25<=315,D30>=0.000000001,D30<=10),0.56*D30^0.25,AND(D25>=315,D25<=400,D30>=0.0000000000001,D30<=0.00000000001),0.0003,AND(D25>=315,D25<=400,D30>=0.00000000001,D30<=0.000000001),0.001,AND(D25>=315,D25<=400,D30>=0.000000001,D30<=10),0.56*D30^0.25,AND(D25>=400,D25<=700,D30>=0.0000000000001,D30<=0.00000000001),0.0000001,AND(D25>=400,D25<=700,D30>=0.00000000001,D30<=0.000005),0.0000002,AND(D25>=400,D25<=700,D30>0.000005,D30<=10),0.003*1.8*D30^0.75,AND(D25>=400,D25<=450,D30>=10,D30<=100),0.01,AND(D25>=450,D25<=500,D30>=10*10^(0.02*(D25-450)),D30<=100),0.001*10^(0.02*(D25-450)),AND(D25>=700,D25<=1050,D30>=0.0000000000001,D30<=0.00000000001),0.0000001,AND(D25>=700,D25<=1050,D30>=0.00000000001,D30<=0.000005),2*10^-7*10^(0.002*(D25-700)),AND(D25>=700,D25<=1050,D30>0.000005,D30<=10),1.8*0.003*D30^0.75*10^(0.002*(D25-700)),AND(D25>=1050,D25<=1150,D30>=0.0000000000001,D30<=0.00000000001),1*10^-7,AND(D25>=1150,D25<=1200,D30>=0.0000000000001,D30<=0.00000000001),10^(0.018*(D25-1150))*10^-7,AND(D25>=1050,D25<=1150,D30>=0.00000000001,D30<=0.000013),2*10^-6,AND(D25>=1150,D25<=1200,D30>=0.00000000001,D30<=0.000013),2*10^(0.018*(D25-1150))*10^-6,AND(D25>=1050,D25<=1150,D30>0.000013,D30<=10),9*D30^0.75*0.001,AND(D25>=1150,D25<=1200,D30>0.000013,D30<=10),9*10^(0.018*(D25-1150))*0.003,AND(D25>=1200,D25<=1400,D30>=0.0000000000001,D30<=0.00000000001),(8+10^(0.04*(D25-1250)))*10^-7,AND(D25>=1200,D25<=1400,D30>=0.00000000001,D30<=0.000013),2*(8+10^(0.04*(D25-1250)))*10^-6,AND(D25>=1200,D25<=1400,D30>0.000013,D30<=10),9*(8+10^(0.04*(D25-1250)))*D30^0.75*0.001,AND(D25>=1400,D25<=1500,D30>=0.0000000000001,D30<=0.00000000001),0.03,AND(D25>=1400,D25<=1500,D30>=0.00000000001,D30<=0.000000001),0.1,AND(D25>=1400,D25<=1500,D30>=0.000000001,D30<=0.001),0.3,AND(D25>=1400,D25<=1500,D30>=0.001,D30<=4),0.56*D30^0.25+0.2,AND(D25>=1400,D25<=1500,D30>=4,D30<=10),1,AND(D25>=1500,D25<=1800,D30>=0.0000000000001,D30<=0.00000000001),0.1,AND(D25>=1500,D25<=1800,D30>=0.00000000001,D30<=0.000000001),0.3,AND(D25>=1500,D25<=1800,D30>=0.000000001,D30<=10),1,AND(D25>=1800,D25<=2600,D30>=0.0000000000001,D30<=0.00000000001),0.01,AND(D25>=1800,D25<=2600,D30>=0.00000000001,D30<=0.000000001),0.03,AND(D25>=1800,D25<=2600,D30>=0.000000001,D30<=0.001),0.1,AND(D25>=1800,D25<=2600,D30>=0.001,D30<=10),0.56*D30^0.25,AND(D25>=2600,D25<=1000000,D30>=0.0000000000001,D30<=0.00000000001),0.001,AND(D25>=2600,D25<=1000000,D30>=0.00000000001,D30<=0.000000001),0.003,AND(D25>=2600,D25<=1000000,D30>=0.000000001,D30<=0.0001),0.01,AND(D25>=2600,D25<=1000000,D30>=0.0001,D30<=10),0.56*D30^0.25,TRUE,\"n/a\")*D31,\"n/a\"),\"n/a\"))","v":"n/a"}]
[null,"MPE for Average Power Limit  (Rule 2)","(J/cm2)",{"$f":"IF(AND(ISBLANK(D24), ISBLANK(D25), ISBLANK(D29)),\n\"n/a\",\nIF(D24=\"Repetitive Pulse\",\n  _xlfn.IFS(\n    AND(D25>=180,D25<=302.5),(3*0.001)/10 * D29,\n    AND(D25>=302.5,D25<=315),(10^(0.2*(D25-295))*0.0004/10) * D29,\n    AND(D25>=315,D25<=400),1/10 * D29,\n    AND(D25>=400,D25<=700),(1*10^-2)/10 * D29,\n    AND(D25>=700,D25<=1050),10^(0.002*(D25-700))*0.001 * D29,\n    AND(D25>=1050,D25<=1150),5*1*0.001 * D29,\n    AND(D25>=1150,D25<=1200),5*10^(0.018*(D25-1150))*0.001 * D29,\n    AND(D25>=1200,D25<=1400),((0.03*(10^(0.01*(1400-D25))))+0.07) * D29,\n    AND(D25>=1400,D25<=1000000),0.1 * D29,\n    TRUE, \"n/a\"),\n\"n/a\"))","v":"n/a"}]
[null,"MPE for Repetitive Pulse Limit (Pulse Train) (Rule 3)","(J/cm2)",{"$f":"IF(AND(ISBLANK(D24), ISBLANK(D38), ISBLANK(D36)),\n\"n/a\",\nIF(D24=\"Repetitive Pulse\",\n   IF(OR(ISBLANK(D38), ISBLANK(D36), NOT(ISNUMBER(D36)), D36<=0), \"n/a\", D38*(D36^-0.025)),\n\"n/a\"))","v":"n/a"}]
[null,"Most Restrictive MPE (Rule 1,2&3)","(J/cm2)",{"$f":"IF(AND(ISBLANK(D24), ISBLANK(D38), ISBLANK(D39), ISBLANK(D40)),\n\"n/a\",\nIF(D24=\"Repetitive Pulse\",\n   IF(COUNTA(D38:D40)=0, \"n/a\", MIN(D38:D40)),\n\"n/a\"))","v":"n/a"}]
[null,"Total Energy Delivered per unit area","(J/cm2)",{"$f":"IF(AND(ISBLANK(D24),ISBLANK(D37),ISBLANK(D36),ISBLANK(D27)),\"\",IF(D24=\"Repetitive Pulse\",IF(AND(ISNUMBER(D37),ISNUMBER(D36),ISNUMBER(D27)),(D37*D36)/(PI()*((D27/10)/2)^2),\"\"),\"n/a\"))","v":"n/a"}]
[null,"Occular MPE for CW or Single Pulsed"," (CW in W/cm2 or Pulsed in J/cm2)",{"$f":"IF(AND(ISBLANK(D24), ISBLANK(D25)),\n\"\",\nIF(D24=\"Continuous\",\n  _xlfn.IFS(\n    AND(D25>=180,D25<=302.5),(3*0.001)/10,\n    AND(D25>=302.5,D25<=315),(10^(0.2*(D25-295))*0.0004/10),\n    AND(D25>=315,D25<=400),1/10,\n    AND(D25>=400,D25<=700),(1*10^-2)/10,\n    AND(D25>=700,D25<=1050),10^(0.002*(D25-700))*0.001,\n    AND(D25>=1050,D25<=1150),5*1*0.001,\n    AND(D25>=1150,D25<=1200),5*10^(0.018*(D25-1150))*0.001,\n    AND(D25>=1200,D25<=1400),((0.03*(10^(0.01*(1400-D25))))+0.07),\n    AND(D25>=1400,D25<=1000000),0.1,\n    TRUE, \"\"),\nIF(D24=\"Pulsed\",\n  _xlfn.IFS(\n    AND(D25>=180,D25<=260,D30>=0.0000000000001,D30<=0.00000000001),0.0003,\n    AND(D25>=180,D25<=260,D30>=0.00000000001,D30<=0.000000001),0.001,\n    AND(D25>=180,D25<=260,D30>=0.000000001,D30<=10),0.56*D30^0.25,\n    AND(D25>=260,D25<=302.5,D30>=0.0000000000001,D30<=0.00000000001),0.0003,\n    AND(D25>260,D25<=302.5,D30>=0.00000000001,D30<=0.000000001),0.001,\n    AND(D25>=260,D25<=302.5,D30>=0.000000001,D30<=10),0.56*D30^0.25,\n    AND(D25>=302.5,D25<=315,D30>=0.0000000000001,D30<=0.00000000001),0.0003,\n    AND(D25>=302.5,D25<=315,D30>=0.00000000001,D30<=0.000000001),0.001,\n    AND(D25>=302.5,D25<=315,D30>=0.000000001,D30<=10),0.56*D30^0.25,\n    AND(D25>=315,D25<=400,D30>=0.0000000000001,D30<=0.00000000001),0.0003,\n    AND(D25>=315,D25<=400,D30>=0.00000000001,D30<=0.000000001),0.001,\n    AND(D25>=315,D25<=400,D30>=0.000000001,D30<=10),0.56*D30^0.25,\n    AND(D25>=400,D25<=700,D30>=0.0000000000001,D30<=0.00000000001),0.0000001,\n    AND(D25>=400,D25<=700,D30>=0.00000000001,D30<=5*0.000001),0.0000002,\n    AND(D25>=400,D25<=700,D30>5*10^-6,D30<=10),0.003*1.8*D30^0.75,\n    AND(D25>=400,D25<=450,D30>=10,D30<=100),1*10^-2,\n    AND(D25>=450,D25<=500,D30>=10*10^(0.02*(D25-450)),D30<=100),0.001*10^(0.02*(D25-450)),\n    AND(D25>=700,D25<=1050,D30>=0.0000000000001,D30<=0.00000000001),0.0000001,\n    AND(D25>=700,D25<=1050,D30>=0.00000000001,D30<=5*10^-6),2*(10^-7)*10^(0.002*(D25-700)),\n    AND(D25>=700,D25<=1050,D30>5*10^-6,D30<=10),1.8*0.003*(D30^0.75)*10^(0.002*(D25-700)),\n    AND(D25>=1050,D25<=1150,D30>=0.0000000000001,D30<=0.00000000001),1*1*10^-7,\n    AND(D25>=1150,D25<=1200,D30>=0.0000000000001,D30<=0.00000000001),1*(10^(0.018*(D25-1150)))*10^-7,\n    AND(D25>=1050,D25<=1150,D30>=0.00000000001,D30<=13*10^-6),2*1*10^-6,\n    AND(D25>=1150,D25<=1200,D30>=0.00000000001,D30<=13*10^-6),2*(10^(0.018*(D25-1150)))*10^-6,\n    AND(D25>=1050,D25<=1150,D30>13*10^-6,D30<=10),9*1*(D30^0.75)*0.001,\n    AND(D25>=1150,D25<=1200,D30>13*10^-6,D30<=10),9*(10^0.018*(D25-1150))*0.003,\n    AND(D25>=1200,D25<=1400,D30>=0.0000000000001,D30<=0.00000000001),1*(8+10^(0.04*(D25-1250)))*1*10^-7,\n    AND(D25>=1200,D25<=1400,D30>=0.00000000001,D30<=13*10^-6),2*(8+10^(0.04*(D25-1250)))*1*10^-6,\n    AND(D25>=1200,D25<=1400,D30>13*10^-6,D30<=10),9*(8+10^(0.04*(D25-1250)))*(D30^0.75)*1*10^-3,\n    AND(D25>=1400,D25<=1500,D30>=0.0000000000001,D30<=0.00000000001),0.03,\n    AND(D25>=1400,D25<=1500,D30>=0.00000000001,D30<=0.000000001),0.1,\n    AND(D25>=1400,D25<=1500,D30>=0.000000001,D30<=0.001),0.3,\n    AND(D25>=1400,D25<=1500,D30>=0.001,D30<=4),(0.56*D30^0.25)+0.2,\n    AND(D25>=1400,D25<=1500,D30>=4,D30<=10),1,\n    AND(D25>=1500,D25<=1800,D30>=0.0000000000001,D30<=0.00000000001),0.1,\n    AND(D25>=1500,D25<=1800,D30>=0.00000000001,D30<=0.000000001),0.3,\n    AND(D25>=1500,D25<=1800,D30>=0.000000001,D30<=10),1,\n    AND(D25>=1800,D25<=2600,D30>=0.0000000000001,D30<=0.00000000001),0.01,\n    AND(D25>=1800,D25<=2600,D30>=0.00000000001,D30<=0.000000001),0.03,\n    AND(D25>=1800,D25<=2600,D30>=0.000000001,D30<=0.001),0.1,\n    AND(D25>=1800,D25<=2600,D30>=0.001,D30<=10),0.56*D30^0.25,\n    AND(D25>=2600,D25<=1000000,D30>=0.0000000000001,D30<=0.00000000001),0.001,\n    AND(D25>=2600,D25<=1000000,D30>=0.00000000001,D30<=0.000000001),0.003,\n    AND(D25>=2600,D25<=1000000,D30>=0.000000001,D30<=0.0000001),0.01,\n    AND(D25>=2600,D25<=1000000,D30>=0.0000001,D30<=10),0.56*D30^0.25,\n    TRUE, \"\"),\n\"n/a\")))","v":""}]
[]
[null,"Nominal Hazard Zone",null,"NHZ, (m)",{"$f":"IF(\n  AND(D24=\"\", D26=\"\", D43=\"\", D41=\"\", D30=\"\"),\n  \"\",\n  IF(D24=\"Continuous\",\n    IF(OR(D43=0, NOT(ISNUMBER(D26)), NOT(ISNUMBER(D43))), \"\", SQRT(D26*COS(0)/(PI()*D43))/100),\n    IF(D24=\"Pulsed\",\n      IF(OR(D43=0, NOT(ISNUMBER(D26)), NOT(ISNUMBER(D43))), \"\", SQRT(D26*COS(0)/(PI()*(D43/D30)))/100),\n      IF(D24=\"Repetitive Pulse\",\n        IF(OR(D30=0, NOT(ISNUMBER(D26)), NOT(ISNUMBER(D41)), NOT(ISNUMBER(D30))), \"\", SQRT(D26*COS(0)/(PI()*(D41/D29)))/100),\n        \"\"\n      )\n    )\n  )\n)","v":""}]
[null,"Nominal Optical Hazard Distance",null,"NOHD, (m) based on 1/e",{"$f":"IF(\n  AND(ISBLANK(D24),ISBLANK(D26),ISBLANK(D27),ISBLANK(D28),ISBLANK(D41),ISBLANK(D43)),\n  \"\",\n  IF(D24=\"continuous\",\n    IF(OR(D28=0,D27=0,D43=0,NOT(ISNUMBER(D26)),NOT(ISNUMBER(D27)),NOT(ISNUMBER(D28)),NOT(ISNUMBER(D43))),\"\",\n      (1000/D28)*(SQRT((4*D26/(PI()*D43*100^2))-(D27/1000)^2))),\n    IF(D24=\"Pulsed\",\n      IF(OR(D28=0,D27=0,D43=0,NOT(ISNUMBER(D26)),NOT(ISNUMBER(D27)),NOT(ISNUMBER(D28)),NOT(ISNUMBER(D43))),\"\",\n        (1000/D28)*(SQRT((4*D26/(PI()*(D43/D30)*100^2))-(D27/1000)^2))),\n      IF(D24=\"Repetitive Pulse\",\n        IF(OR(D28=0,D27=0,D41=0,NOT(ISNUMBER(D26)),NOT(ISNUMBER(D27)),NOT(ISNUMBER(D28)),NOT(ISNUMBER(D41))),\"\",\n         (1000/D28)*(SQRT((4*D26/(PI()*(D41/D29)*100^2))-(D27/1000)^2))),\n        \"\"\n      )\n    )\n  )\n)","v":""}]
[null,"Optical Density of Eyewear Required",null,"OD",{"$f":"IF(OR(ISBLANK(D24),ISBLANK(D35),ISBLANK(D43),ISBLANK(D37),ISBLANK(D31),ISBLANK(D30),ISBLANK(D41)),\"\",ROUNDUP(IF(D24=\"Continuous\",LOG10(D35/D43),IF(D24=\"Pulsed\",LOG10(D35/(D43/D30)),IF(D24=\"Repetitive Pulse\",LOG10((D35/(D41/D29))),0))),0))","v":""}]
[]
[null,"Interlock and/or Warning Light",null,{"$f":"IF(ISBLANK(D22),\"\",IF(D22=4,\"Required\",\"Recommended\"))","v":""},"Laser Warning Signages Template (choose applicable below)"]
[null,"Warning Signs at Entry Ways",null,{"$f":"IF(ISBLANK(D22),\"\",IF(D22=4,\"Warning\", \"Warning\"))","v":""},"Laser signage for labs with both - class 3b and 4"]
[null,"Curtains and Barriers",null,{"$f":"IF(\n  D24=\"\",\n  \"\",\n  IF(D24=\"Continuous\",\n    IF(OR(D35=\"\", D43=\"\"), \"\", IF(D35>D43, \"Required\", \"Not Required\")),\n  IF(D24=\"Pulsed\",\n    IF(OR(D35=\"\", D43=\"\"), \"\", IF(D35>D43, \"Required\", \"Not Required\")),\n  IF(D24=\"Repetitive Pulse\",\n    IF(OR(D35=\"\", D41=\"\"), \"\", IF(D35>D41, \"Required\", \"Not Required\")),\n  \"\"\n))))","v":""},"Laser signage for class - 3b"]
[null,"Ignition Hazard (>0.5W/cm2)",null,{"$f":"IF(TRIM(D35)=\"\",\"\",IF(D35>=0.5, \"Yes\", \"No\"))","v":""},"Laser signage class - 4"]
[]
[null,"IV. SOP and Safety Checklist"]
[]
[null,"Note: Confirm if parameter is applicable to current set-up and laser area and attest affirmation of compliance (Justify if you choose \"Not Comply\" under Remarks below the form). Please use the drop-down selections provided."]
[null,"POTENTIAL HAZARDS",null,null,"Applicable/ Not Applicable","Statement of Compliance"]
[null,"Laser Hazards","Direct beam and reflections can cause eye damage (blindness) and skin injury"]
[null,null,"Diffuse reflections also pose eye risks"]
[null,null,"Others (Please Specify):"]
[null,"Electrical Hazards:","Risk of shock/electrocution from electrical sources and energized conductors"]
[null,null,"Ensure hands/gloves are dry before plugging in power cords"]
[null,null,"Others (Please Specify):"]
[null,"Chemical Hazards:","Keep flammable solvents out of the laser beam path"]
[null,null,"Others (Please Specify):"]
[null,"Others (Please Specify):"]
[null,"WORK PRACTICE CONTROLS",null,null,"Applicable/ Not Applicable","Statement of Compliance"]
[null,"Lasers Controls:","Only authorized users may operate lasers"]
[null,null,"Keep doors closed and warning lights ON during use and alignment"]
[null,null,"Supervise unauthorized personnel with PI approval"]
[null,null,"Wear appropriate eyewear for laser (available on a on wall rack)"]
[null,null,"Always wear eyewear when laser is on"]
[null,null,"Use beam blocks, housings, and enclosures to control reflections"]
[null,null,"Follow alignment steps in System Alignment Section"]
[null,null,"Survey and document stray beams; notify users of changes"]
[null,null,"Use same safety steps when adjusting optics"]
[null,null,"New users must do a beam survey before use"]
[null,null,"Others (Please Specify):"]
[null,"Electrical Controls:","Unplug power before removing enclosures; follow manufacturer safety steps"]
[null,null,"Only qualified personnel may do maintenance; two-person rule applies"]
[null,null,"Treat all components as live"]
[null,null,"Others (Please Specify):"]
[null,"Chemical Controls:","Check flash point and temperature rise before using flammable solvents under the beam"]
[null,null,"Others (Please Specify):"]
[null,"Others (Please Specify):"]
[null,"PERSONAL PROTECTIVE EQUIPMENT (PPE)",null,null,"Applicable/ Not Applicable","Statement of Compliance"]
[null,"PPEs","Provide appropirate Laser Eyewear  "]
[null,null,"UV-blocking clothing for high-intensity UV lasers"]
[null,null,"Long pants and closed-toe shoes required"]
[null,null,"Others (Please Specify):"]
[null,"Others (Please Specify):"]
[null,"EXPERIMENTAL PROCEDURE",null,null,"Applicable/ Not Applicable","Statement of Compliance"]
[null,"A. Normal Operation","Check electrical and water connections"]
[null,null,"Remove reflective jewelry/clothing"]
[null,null,"Get and wear correct OD eyewear"]
[null,null,"Turn on warning light and close doors/curtains"]
[null,null,"Remove optical setup cover and inspect for changes"]
[null,null,"Ensure beam path is blocked and controls are in place"]
[null,null,"Confirm everyone is wearing eyewear"]
[null,null,"Give verbal warning, then:"]
[null,null,"Remove laser cap"]
[null,null,"Turn on power supply"]
[null,null,"Insert key and power on laser"]
[null,null,"Others (Please Specify):"]
[null,"B. System Alignment","Use low-power laser if possible"]
[null,null,"Complete Normal Operation checklist"]
[null,null,"If needed, reduce beam to lowest power"]
[null,null,"Align beam using viewing card/electronic viewer"]
[null,null,"Block stray beams and secure all mounts"]
[null,null,"Announce power increase if others are present"]
[null,null,"Follow safety tips:"]
[null,null,"Only trained personnel present"]
[null,null,"Minimize people in lab"]
[null,null,"Avoid eye-level beam paths"]
[null,null,"Use visible lasers for simulation"]
[null,null,"Block reflections close to source"]
[null,null,"Use beam shutters when not in use"]
[null,null,"Others (Please Specify):"]
[null,"C. System Shutdown","Turn off laser and remove key"]
[null,null,"Turn off power supply"]
[null,null,"Cap laser and cover optics"]
[null,null,"Return eyewear to rack"]
[null,null,"Turn off warning light"]
[null,null,"Others (Please Specify):"]
[null,"D. Emergency Procedures","Shut down laser and ensure personnel safety"]
[null,null,"Seek medical help:"]
[null,null,"UCF employees: 407-266-3627"]
[null,null,"Students: 407-823-2701"]
[null,null,"After hours: Advent Health ER – 407-281-3600"]
[null,null,"Notify PI and file Incident Report"]
[null,null,"Contact Laser Safety Officer: 407-823-0476 or Mario.DeVera@ucf.edu"]
[null,null,"Others (Please Specify):"]
[null,"E. Others (Please Specify):"]
[null,"TRAINING",null,null,"Applicable/ Not Applicable","Statement of Compliance"]
[null,"A. Authorized Personnel","Only authorized users may operate the laser"]
[null,null,"Must follow all safety procedures outlined in the SOP"]
[null,null,"Others (Please Specify):"]
[null,"B. Steps to Become Authorized","Complete EHS 309 online laser safety training"]
[null,null,"(Optional) Get a baseline eye exam per ANSI Z136.1"]
[null,null,"Read and understand the SOP"]
[null,null,"Receive hands-on training from an authorized user"]
[null,null,"Sign the authorized user sheet to confirm completion"]
[null,null,"Others (Please Specify):"]
[null,"C. Unauthorized Personnel","May enter only with supervision during laser operation"]
[null,null,"Must be briefed on safety protocols"]
[null,null,"Must wear appropriate laser eyewear"]
[null,null,"Others (Please Specify):"]
[null,"REMOVAL FROM INVENTORY",null,null,"Applicable/ Not Applicable","Statement of Compliance"]
[null,"Removal from   Inventory","Report to EHS Serial no. of laser for disposal or relocation/transfer"]
[null,null,"Through EHS, verify if removal of laser from FDOH inventory listings "]
[null,null,"Accomplish Notice of Decontamination, if applicable."]
[null,null,"Others (Please Specify):"]
[]
[null,"V. Laser Layout"]
[]
[null,"Please sketch or draw  the laser layout in the lab/room using shapes available in the toolbar. If available,  insert picture of laser layout by going to Insert  → Pictures in the toolbar to add you image inside the box below."]
[]
[null,"Laser Layout"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null,"Remarks:"]
[]
[]
[null,"PI Name and Signature:"]
[null,"Date:"]
[]
[]
[null,"(Submit accomplished form to Mario.DeVera@ucf.edu) "]
[]
[null,null,null,null,"Revision 1.3","2025-08-29"]

{"sheet":"Dropdowns","index":1,"n_rows":7,"n_cols":2,"rows_kept":7,"cut":false,"hidden":true,"merged":[]}
[]
["Please Select","Applicable "]
["Continuous","Not Applicable"]
["Pulsed"]
["Repetitive Pulse","Comply "]
[null,"Not Comply"]
[null,"N/A"]